## Slide 1: An Adaptive Calibration Algorithm Using Progressive Measurement and Accelerated Measurement Prediction
Yuxuan Chen

## Slide 2: Traditional Method
IC
PC
Computer sends all settings to the DUT (in arbitrary order) for measurement.
A constellation of all settings is generated.
Computer scans all constellation points to find the point closest to the ideal position and establish the mapping.
Advantage:
Highest accuracy and often guaranteed solution.
Measurement order and details of the controls are hidden from the user.
Relaxed requirement for the DUT gain and phase response.
Disadvantage:
Long measurement time and long post-processing time. (Need to measure every point and computer need to scan every point for comparison for every ideal points)

## Slide 3: Adaptive Measurements

## Slide 4: Adaptive Measurements

## Slide 5: Adaptive Measurements
Need bi-directional relationship from DUT settings to the measurement points.
DUT response need to be predictable, and the measurement results need to be near the ideal points.

## Slide 6: Outphasing Phase Shifter

## Slide 7: Outphasing Phase Shifter

## Slide 8: Outphasing Phase Shifter

## Slide 9: Outphasing Phase Shifter

## Slide 10: Outphasing Phase Shifter

## Slide 11: Outphasing Phase Shifter

## Slide 12: Outphasing Phase Shifter

## Slide 13: Outphasing Phase Shifter

## Slide 14: Outphasing Phase Shifter

## Slide 15: Outphasing Phase Shifter
Vector Phases to Cartesian:
Cartesian to Vector Phases:
(using inverse kinematic)

## Slide 16: Outphasing Phase Shifter
Vector Phases to Cartesian:
Cartesian to Vector Phases:
(using inverse kinematic)

## Slide 17: RFVM
Vector lengths to Cartesian:
Cartesian to Vector lengths:

## Slide 18: Adaptive Measurements
ideal target point
For an ideal target point, use inverse kinematic to calculate the corresponding two vector phases and obtain the measurements.

## Slide 19: Adaptive Measurements
ideal target point
For an ideal target point, use inverse kinematic to calculate the corresponding two vector phases and obtain the measurements.
measurements
If measure all the points for a constant gain circle at once, the distorted circle makes adjustment difficult and reductant measurements are made.

## Slide 20: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0

## Slide 21: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.

## Slide 22: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.

## Slide 23: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.

## Slide 24: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.

## Slide 25: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.

## Slide 26: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.

## Slide 27: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.

## Slide 28: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.

## Slide 29: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution)

## Slide 30: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution)

## Slide 31: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.

## Slide 32: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.

## Slide 33: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
  is targeting point inside the measurement results boundary?

## Slide 34: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.

## Slide 35: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.

## Slide 36: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.

## Slide 37: Point-to-Point Progressive Measurements

## Slide 38: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.

## Slide 39: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.

## Slide 40: RTPS Phase Switching Modes
RTPS 1
RTPS 2
RTPS 3
RTPS 1
RTPS 2
RTPS 3

## Slide 41: RTPS Phase Switching Modes

## Slide 42: RTPS Phase Switching Modes

## Slide 43: RTPS Phase Switching Modes

## Slide 44: RTPS Phase Switching Modes
RTPS 1
RTPS 2
RTPS 3
RTPS 1
RTPS 2
RTPS 3
9 Modes for a pair of phase setting.
Due to PVT variation and mismatch, the constellation of the 9 modes are slightly different.
Provide fine tune to the constellation point.
Due to the random nature of this variation, this effect cannot be learned.

## Slide 45: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.

## Slide 46: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.

## Slide 47: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.

## Slide 48: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.
  Measure the 9 modes and find the closest point to the ideal point.

## Slide 49: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.
  Measure the 9 modes and find the closest point to the ideal point.

## Slide 50: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.
  Measure the 9 modes and find the closest point to the ideal point.
  Establish mapping between the codeword of the closest point to the position of the ideal point.

## Slide 51: Point-to-Point Progressive Measurements
Phase offset compensation
  Target an ideal point of phase 0
  Iteratively adjust phase offset so that the measurement point aligns with the target point.
Next point calibration
  Increment the ideal point by the target phase resolution.
  Perform measurements in the search area (± actual phase resolution and ± actual gain resolution).
  Find the measurement point in the search area that is closet to the ideal point.
  In the case of deviating measurements, a test is performed to evaluate the search area measurement validity.
    The error vectors for every ideal/measurement pair are averaged and a new search area is found in the opposite direction.
  Measure the 9 modes and find the closest point to the ideal point.
  Establish mapping between the codeword of the closest point to the position of the ideal point.
  Repeat 2 until all ideal points on this gain circle are calibrated.

## Slide 52: Point-to-Point Progressive Measurements
Remarks:
Progressive measurement is repeated for every needed constant gain circle.
The step of phase offset compensation should be done after DC offset compensation and vector length measurement.
All measurement results in the search area are saved so that there is no need for repetitive measurement when accessed again.

## Slide 53: Accelerated Measurement Prediction

## Slide 54: Accelerated Measurement Prediction

## Slide 55: Accelerated Measurement Prediction

## Slide 56: Accelerated Measurement Prediction

## Slide 57: Accelerated Measurement Prediction
Mapping

## Slide 58: Accelerated Measurement Prediction

## Slide 59: Accelerated Measurement Prediction

## Slide 60: Accelerated Measurement Prediction

## Slide 61: Accelerated Measurement Prediction

## Slide 62: Accelerated Measurement Prediction

## Slide 63: Accelerated Measurement Prediction

## Slide 64: Accelerated Measurement Prediction

## Slide 65: Accelerated Measurement Prediction

## Slide 66: Accelerated Measurement Prediction
Target Position
Ideal Position that results in the closest measurement to the target position

## Slide 67: Accelerated Measurement Prediction
Target Position
Ideal Position that results in the closest measurement to the target position
In-situ training and validating

## Slide 68: Accelerated Measurement Prediction

## Slide 69: Accelerated Measurement Prediction

## Slide 70: Accelerated Measurement Prediction

## Slide 71: Accelerated Measurement Prediction

## Slide 72: Accelerated Measurement Prediction
Target Position
Ideal Position that results in the closest measurement to the target position
In-situ training and validating
Progressive training per gain circle.
Neural Network replaces the normal routine after making 5 correct prediction in a row.
When the network is not validated, a new targeting point is fed to the neural network for testing before it is used for training so that the testing data has not been trained by the network.
Neural Network prediction is verified by doing the 9 mode measurements and judge if the closet point is close enough (by previous standard).
If Neural Network prediction is checked to be inaccurate, the current targeting point is calibrated using the normal routine and the neural network restart training.

## Slide 74: Thank you

## Slide 75: An Adaptive Calibration Algorithm Using Progressive Measurement and Accelerated Measurement Prediction
Yuxuan Chen
Development Update 1

## Slide 76: Demo

## Slide 77
Calibration Result (Single Mode)

## Slide 78
Traditional Method (Single Mode)

## Slide 79
Measurement non-uniformity causes distortion in constant gain circle

## Slide 80
Measurement non-uniformity causes distortion in constant gain circle
Causes the measurement points to go to unexpected locations.
Causes measurement dead zones.
Distortion gets weaker as the gain circle radius increases.
Expect to see RMS errors gets better for larger gain circles. (Verified)

## Slide 81
Measurement Dead zones

## Slide 82
Measurement Dead zones
Decrease the gain and phase steps in the search kernel (to below the theoretical resolutions).
Use absolute value instead of dB for gain resolution.
Implement adaptive gain resolution (observe no improvement).
For each ideal point, search for the two inverse kinematic solutions which doubles the number of measurements.

## Slide 83
Measurement Dead zones

## Slide 84
Measurement Dead zones

## Slide 85
Crossing the sea
Adaptively increase the search area kernel size if the program is stuck on calibrating an ideal point for too long.
Filter the measurement results and remove outliers for kernel validation.

## Slide 86: An Adaptive Calibration Algorithm Using Progressive Measurement and Accelerated Measurement Prediction
Yuxuan Chen
Development Update 2

## Slide 87
Calibration for individual gain and phase control chain
8-bit VGA and 8-bit phase shifter
256*256 = 65536 points
Phase corrected to reference the phase of “0” phase setting for each gain settings.

## Slide 88
Gain Profile Characterization
For better initial guess of the gain index for a targeting gain.
Gain profile is scaled by the max gain ratio in the targeting phase direction.

## Slide 89
Chart: 
| Kernel Search | Traditional
Column3 | 5593 | 65536
Number of Measurement

## Slide 90
Chart: 
|  |  | Kernel Search | Traditional
Column3 | 783 | 1170 | 5593 | 65536
Number of Measurement

## Slide 91: Demo

## Slide 92
Chart: 
| Machine Learning | Optimization | Kernel Search | Traditional
Column3 | 783 | 1170 | 5593 | 65536
Number of Measurement

## Slide 93
For calibrating 7 gain circles of 64 points each
Traditional:
  65536 points measured
  146.3 measurement per point
Kernel Search:
  5593 points measured (275 for pre-characterization)
  12.5 measurement per point (11.9 excluding the pre-characterization)
Single-Point Optimization + Kernel Search:
  1170 points measured (275 for pre-characterization)
  2.6 measurement per point (2 excluding the pre-characterization)
Machine Learning + Single-Point Optimization + Kernel Search:
  783 points measured (275 for pre-characterization)
  1.75 measurement per point (1.13 excluding the pre-characterization)

## Slide 94
Single-Point Optimization
For well-behaved variable-gain phase shifter and relaxed RMSE requirement, single-point optimization method often arrive to an acceptable point in less step than the measurement needed for a kernel search.
If optimization cannot find an acceptable point within certain steps, then the kernel search activates.

## Slide 95
Machine Learning for accurate first guess
A baseline model is trained with a vast amount of measurement data of a single PS type.
  Normalized gain and phase-corrected and normalized phase.
  Gain from 0 to 1. Phase from 0 to 1.
The baseline model continue to incrementally train using the new measurement data for adaption.

## Slide 96: An Adaptive Calibration Algorithm Using Progressive Measurement and Accelerated Measurement Prediction
Yuxuan Chen
Development Update 3

## Slide 97
Vector Sum
Two 6-bit VGA
64*64 = 4096 points
Constellation not rotated in consideration of constellation rotation mismatch in array level calibration.

## Slide 98
For calibrating 7 gain circles of 64 points each
Traditional:
  65536 points measured
  146.3 measurement per point
Kernel Search:
  5593 points measured (275 for pre-characterization)
  12.5 measurement per point (11.9 excluding the pre-characterization)
Single-Point Optimization + Kernel Search:
  1170 points measured (275 for pre-characterization)
  2.6 measurement per point (2 excluding the pre-characterization)
Machine Learning + Single-Point Optimization + Kernel Search:
  783 points measured (275 for pre-characterization)
  1.75 measurement per point (1.13 excluding the pre-characterization)

## Slide 99
For calibrating 7 gain circles of 64 points each
1.75 1.13
Total Excluding Pre-characterization

## Slide 100
For calibrating 7 gain circles of 64 points each
1.24 0.99
Total Excluding Pre-characterization

## Slide 101: Demo

## Slide 102
Total Phase RMSE: 1.102 degrees
Total Gain RMSE: -23.779 dB
This Algorithm: 556 measurements
Total Phase RMSE: 0.884 degrees
Total Gain RMSE: -24.201 dB
Traditional: 4096 measurements

## Slide 103
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 104
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 105
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 106
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 107
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 108
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 109
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 110
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 111
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 112
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 113
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 114
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 115
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 116
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 117
Vector Sum Pre-characterization
Need a reference point as the origin of the two vectors
Need the direction of the two vectors
Need a profile to predict the index/code needed to get to a specific vector length

## Slide 118
Vector Sum Pre-characterization
Reference point doesn’t need to be at DC offset > No need for measuring DC offset
The vector profile near the reference point need to be well defined > select reference point near the center of the constellation but it doesn’t have to be THE CENTER.

## Slide 121: An Adaptive Calibration Algorithm Using Progressive Measurement and Accelerated Measurement Prediction
Yuxuan Chen
Phase One Summary

## Slide 122
Formulation

## Slide 125
Outphasing

## Slide 126
Outphasing

## Slide 127
Outphasing

## Slide 128
Separate Indexing

## Slide 129
Separate Indexing

## Slide 130
Separate Indexing

## Slide 131
Vector Sum

## Slide 132
Vector Sum

## Slide 133
Vector Sum

## Slide 134
Calibration Mechanism
Ideal Point Position
Point Translation
Code/Index
Measurement
Measurement Point Position
Point Finding
Next*

## Slide 135
Calibration Mechanism
Ideal Point Position
Point Translation
Code/Index
Unique for different types of variable gain phase shifters
Provides a first guess of the code/index
Established through:
  Pre-characterization
    DC offset, vector length, phase offset, index-to-position profile
  Machine Learning

## Slide 136
Calibration Mechanism
Next ideal Point Position
Point Finding
Measurement
Same for different types of variable gain phase shifters (it’s based on cartesian coordinates)
Only needed if the first guess doesn’t satisfy specified requirements
Include:
  Measurement filtering and validation
  Adaptive kernel search
  Kernel space mapping

## Slide 137
Two Types of Solution
  Solution by Encirclement
The closest measurement point is chosen if the ideal point is enclosed by the measurements from the kernel.
Only needed if all measurement points from the kernel fails the vicinity test.
Guarantee local best solution if the constellation mapping is continuous.
Solution by Vicinity
A measurement point is chosen if it is close enough to the ideal position by a specified standard.
Faster and less measurement needed.

## Slide 138
Shape of Constellation
Fully-Filled Circular
Fully-Filled
Distorted Circular
Fully-Filled
Distorted Square

## Slide 139
Shape of Constellation
Partially-Filled Circular
Partially-Filled
Distorted Circular
Fully-Filled
Distorted Square
Lower gain is bounded by lowest gain setting in Separate indexing

## Slide 140
Point Translation

## Slide 141
Point Translation
Outphasing
Separate-Indexing
Vector-Sum

## Slide 142
Point Translation
Outphasing
Separate-Indexing
Vector-Sum

## Slide 143
Outphasing
Cartesian to Vector Phases:
[phase1, phase2] = cartesian2phases(X, Y)

## Slide 144
Outphasing
For accurate first guess:
Need to reference points to DC offset
Need accurate vector lengths L1 and L2
Need to compensate for the two vector’s phase offset

## Slide 145
Outphasing
For accurate first guess:
Need to reference points to DC offset
Need accurate vector lengths L1 and L2
Need to compensate for the two vector’s phase offset

## Slide 146
Outphasing: Pre-Characterization
DC Offset:
Set two vectors to have the same setting and sweep a circle.
[0 0; pi/2 pi/2; pi pi; 3*pi/2 3*pi/2; pi/4 pi/4; 3*pi/4 3*pi/4; 5*pi/4 5*pi/4; 7*pi/4 7*pi/4]

## Slide 147
Outphasing: Pre-Characterization
DC Offset:
Set two vectors to have the same setting and sweep a circle.
Calculate the average of all the measured points
cartesian point from summing vectors of angle i and j (ideal position)
cartesian point from measuring (measurement position)

## Slide 148
Outphasing: Pre-Characterization
DC Offset Compensation:
Subtract the targeting point by the DC offset point.

## Slide 149
Outphasing: Pre-Characterization
Vector Length:
For a constant angle of one vector, sweep the other vector.

## Slide 150
Outphasing: Pre-Characterization
Vector Length:
For a constant angle of one vector, sweep the other vector.
Find the center of each circle by finding the average of the measurements on each circle.
cartesian point of circle center whose angle to the DC offset is i.

## Slide 151
Outphasing: Pre-Characterization
Vector Length:
For a constant angle of one vector, sweep the other vector.
Find the center of each circle by finding the average of the measurements on each circle.
Find average radius of each circle.
average radius of circle with center of

## Slide 152
Outphasing: Pre-Characterization
Vector Length:
For a constant angle of one vector, sweep the other vector.
Find the center of each circle by finding the average of the measurements on each circle.
Find average radius of each circle.
Calculate L1/L2 by averaging radius of average radii.

## Slide 153
Outphasing: Pre-Characterization
Vector Length:
For a constant angle of one vector, sweep the other vector.
Find the center of each circle by finding the average of the measurements on each circle.
Find average radius of each circle.
Calculate magnitude scaling factor (M) for normalizing reading.
Calculate L1/L2 by averaging radius of average radii.

## Slide 154
Outphasing: Pre-Characterization
Phase Offset:
The vector lengths are known, and the target points are DC offset compensated.
The common phase offset will cause rotation of the constellation and the difference phase offset will cause gain error.

## Slide 155
Outphasing: Pre-Characterization
Phase Offset:
The vector lengths are known, and the target points are DC offset compensated.
The common phase offset will cause rotation of the constellation and the difference phase offset will cause gain error.

## Slide 156
Outphasing: Pre-Characterization
Phase Offset:
The vector lengths are known, and the target points are DC offset compensated.
The common phase offset will cause rotation of the constellation and the difference phase offset will cause gain error.

## Slide 157
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.

## Slide 158
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.

## Slide 159
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.

## Slide 160
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 161
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 162
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 163
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 164
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 165
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 166
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 167
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 168
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 169
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 170
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 171
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 172
Outphasing: Pre-Characterization
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phases needed to get to the measured point.
Iteratively compensate for this phase offset until the measured point pass the vicinity test of the ideal point.

## Slide 173
Outphasing: Pre-Characterization
Phase Offset Compensation:
Subtract the targeting phases by their corresponding phase offset.

## Slide 174
Outphasing: Index Mapping
Linear Mapping
is device specific.
Follow 1-to-1 linear mapping between the compensated phases for a phase shifter to its index range.
For two phase shifters each covers 360 degree:

## Slide 175
Point Translation
Outphasing
Separate-Indexing
Vector-Sum

## Slide 176
Point Translation
Outphasing
Separate-Indexing
Vector-Sum

## Slide 177
Point Translation
Outphasing
Separate-Indexing
Vector-Sum

## Slide 178
Separate-Indexing
Gain index
Phase index

## Slide 179
Separate-Indexing
Gain index
Phase index

## Slide 180
Separate-Indexing: Phase Index-Mapping
Gain index: i Phase index: j
cartesian point from measuring gain and phase index of i and j

## Slide 181
Phase Offset:
Phase reference is set to the phase of the measurement injection signal.
j = 0
Separate-Indexing: Phase Index-Mapping

## Slide 182
Phase Offset:
For each gain circle, set the ideal point to be of phase 0.
Measure this point and calculate the phase of the measured point.
Iteratively compensate for this phase offset until the measured phase is smaller than the actual phase resolution.
Separate-Indexing: Phase Index-Mapping

## Slide 183
Phase Offset Compensation:
Subtract the targeting phase by the phase offset.
Separate-Indexing: Phase Index-Mapping

## Slide 184
Separate-Indexing
Need pre-characterization to build an index-to-gain/phase profile to accurately model the relationship between gain/phase and index variation.
For reducing the number of measurement in pre-characterization, the profile needs to be based on general phenomena thus it could be applied in different occasions.
The profile doesn’t have to be precise, but it need to put the measurement point in the vicinity of the kernel search area of the ideal point thus a solution can converge quickly.

## Slide 185
Separate-Indexing: Pre-Characterization
We expect some phase offset and it will be compensated at every targeting gain circle
Sampling Gain index: {round(linspace(1,end,steps))}
Sampling Phase index: {1} (arbitrary)

## Slide 186
Separate-Indexing: Pre-Characterization
Sampling Gain index: {round(linspace(1,end,steps))}
cartesian point from measuring gain and phase index of i and j
Sampling Phase index: {1} (arbitrary)

## Slide 187
Separate-Indexing: Pre-Characterization

## Slide 188
Separate-Indexing: Pre-Characterization

## Slide 189
Separate-Indexing: Pre-Characterization

## Slide 190
Separate-Indexing: Pre-Characterization
Sampling Gain index: {end}
Sampling Phase index: {round(linspace(1,end,steps))}
……

## Slide 191
Separate-Indexing: Pre-Characterization
Sampling Gain index: {end}
Sampling Phase index: {round(linspace(1,end,steps))}
……

## Slide 192
Separate-Indexing: Pre-Characterization
***
The gain profile assumes the lowest gain circle follows the footprint of the largest gain circle.
The larger the phase variation with gain, the less linearly-scalable the gain profile is.
More complicated gain profile translation technique is needed if the phase variation is large and you intent to calibrate for gain circles very close to the lowest available gain.

## Slide 193
Separate-Indexing: Pre-Characterization
***
The gain profile assumes the lowest gain circle follows the footprint of the largest gain circle.
The larger the phase variation with gain, the less linearly-scalable the gain profile is.
More complicated gain profile translation technique is needed if the phase variation is large and you intent to calibrate for gain circles very close to the lowest available gain.
This effect can be ignored if the lowest gain circle is very small.

## Slide 194
Separate-Indexing: Pre-Characterization
Constant phase lines
Constant gain circles
What about phase profile?
Phase profile often changes with phase, which makes it unscalable.

## Slide 195
Separate-Indexing: Pre-Characterization
What about phase profile?
Phase profile often changes with phase, which makes it unscalable.
If maximum phase variation in between gain indexes are less than the targeting phase resolution (small enough for kernel search), then phase profile is unnecessary.

## Slide 196
Separate-Indexing: Pre-Characterization
What about phase profile?
Phase profile often changes with phase, which makes it unscalable.
If maximum phase variation in between gain indexes are less than the targeting phase resolution (small enough for kernel search), then phase profile is unnecessary.
Leave phase variation to kernel search or use machine learning.